TEST METHOD & COMPLIANCE DOSSIER

ANSI/ASHRAE 110

Method of Testing Performance of Laboratory Fume Hoods

A test-method guide and compliance dossier for Deiiang fume hoods

What each test measures, what instrument produces the number, and where the acceptance limits actually come from.

Deiiang Cleanroom World · Laboratory Division   |   Edition 1.0 · October 2026

CONTENTS

1. Purpose and scope

2. The standard in one page

3. The three test states: AM, AI, AU

4. Test 1 — Face velocity

5. Test 2 — Flow visualisation

6. Test 3 — Tracer gas containment

7. Acceptance criteria and the citation trap

8. Deiiang product compliance matrix

9. The control layer: why a test result is reproducible

10. As-installed commissioning: a twelve-step protocol

11. As-used verification and the annual cycle

12. Why compliant hoods still leak

13. What a defensible test report contains

14. Limitations and honest notes

15. References

Appendix A — As-installed commissioning data sheet

Appendix B — Abbreviations

Summary. ASHRAE 110 measures three things and sets no pass mark. Face velocity screens; visualisation explains; tracer gas quantifies. The limits everyone quotes, 0.05 and 0.10 ppm, come from Z9.5, not from ASHRAE 110. This guide sets out the arithmetic, then shows which Deiiang design provision produces which result, what has been measured on the T3 platform, and how to commission a hood so the number still holds in year eight.

| Item | What this guide establishes |
| --- | --- |
| Three tests, one response test | Face velocity, flow visualisation, tracer gas containment, plus the sash movement effect test that exercises the VAV control loop |
| Three test states | AM, AI and AU answer three different questions and are performed by three different parties; AM results are not transferable to a project |
| Control levels | AM 0.05 ppm; AI and AU 0.10 ppm, from Z9.5-2022 §3.4 — a report that omits the source is not auditable |
| Measured evidence | T3 platform: 0.50 m/s mean, 6.0 % uniformity, 41 Pa resistance, 0.00 ppm tracer at three positions, 0.00 ppm RA-Max |

1. Purpose and scope

This section defines what the document covers, what it deliberately leaves out, and which edition of each reference it relies on.

1.1 What this document is

This is an engineering guide to ANSI/ASHRAE 110, Method of Testing Performance of Laboratory Fume Hoods, written for the people who have to specify, commission, accept and periodically re-verify a fume hood rather than for the laboratory technician who simply uses one.

It explains what each of the three tests in the standard actually measures, what instrument produces the number, how the number is reduced to a reported figure, and where the acceptance limits come from — because they do not come from ASHRAE 110.

The second half of the document applies the method to the Deiiang fume hood range: which design feature of a given hood exists in order to produce a given test result, what measured data exists today, and what a project-specific test programme should look like from factory test through annual re-verification.

1.2 What ASHRAE 110 is — and what it is not

ASHRAE 110 is a test method. It prescribes apparatus, positions, release rates, sampling durations and the arithmetic by which raw readings become a reported containment figure. It deliberately does not state a pass/fail threshold, because the acceptable leakage from a hood depends on the toxicity of what is handled inside it and on the exposure judgement of the responsible person.

The thresholds normally quoted in specifications — 0.05 ppm as-manufactured, 0.10 ppm as-installed and as-used — come from ANSI/ASHRAE/ASSP Z9.5, Laboratory Ventilation, and in some markets from SEFA 1 or from the owner's own purchasing specification. A test report that says "passed ASHRAE 110" without naming the control level and its source is not auditable.

Section 7 of this guide deals with that citation trap in detail, because it is the single most common defect in fume hood acceptance documentation.

1.3 Who should read which part

| Reader | Read |
| --- | --- |
| Owner, EHS, chemical hygiene officer | 2, 3, 7, 10, 11, 12 |
| Design institute / consulting engineer | 2 to 8, 10, 13 |
| Procurement and tender evaluation | 3, 7, 8, 13, Appendix A |
| Commissioning and test agency | 4 to 7, 10, 13, Appendix A |
| Distributor and export sales | 1, 2, 8, 9, 12 |

1.4 Normative references used

| Reference | Edition used here | Role in this document |
| --- | --- | --- |
| ANSI/ASHRAE 110 | 2016, with published addenda | The test method itself |
| ANSI/ASHRAE/ASSP Z9.5 | 2022 | Control levels, design face velocity, routine testing |
| SEFA 1 | Current edition | AM testing obligation, cross-draft guidance |
| EN 14175-2 / -3 | 2003 / 2006 | European counterpart, for cross-market comparison |
| JGJ 91-2019 | 2019 | Chinese tiered face velocity values |
| GB/T 24820-2024 | 2024 | Chinese laboratory furniture safety standard |

Where the editions differ, the contract should name the edition. "ASHRAE 110" unqualified is an ambiguous specification.

2. The standard in one page

ASHRAE 110 contains three tests, run in a fixed order, plus one conditional response test. Everything else in the standard is apparatus definition and reporting discipline.

| Test | What it measures | Instrument | Reported as |
| --- | --- | --- | --- |
| Face velocity | Mean and distribution of air speed across the working opening | Calibrated thermal anemometer on a grid | Mean m/s, per-point values, uniformity |
| Flow visualisation | Qualitative containment and airflow pattern | TiCl₄ stick and smoke generator | Pass / fail by observation, clearance time |
| Tracer gas containment | Quantitative leakage to the breathing zone | SF₆ ejector, manikin, gas analyser | 5-min average, minimum, maximum in ppm |
| Sash movement effect | Containment while the sash is moving and the VAV responds | Same as tracer gas, dynamic | 45 s rolling average (RA-Max), peak |

The order matters. Face velocity is measured first because a hood that cannot hold its design airflow cannot be expected to contain anything, and because the measurement is cheap. Tracer gas is measured last because it is expensive, slow, and only meaningful once the airflow is proven stable.

One structural point is worth stating plainly at the outset: the three tests are not correlated. A hood can hold 0.50 m/s across its opening and still leak; the Iranian petrochemical survey of 22 hoods found face velocity compliance in 9.1 % and tracer gas compliance in 4.5 %, with no hood passing one test passing the other. The Finnish survey of 295 hoods found 27 % compliant on face velocity but only 2.7 % compliant on every criterion.

Any acceptance regime that stops at face velocity is measuring the wrong thing.

3. The three test states: AM, AI, AU

ASHRAE 110 defines three states in which the same method may be applied. They are not three grades of the same test; they answer three different questions, and they are performed by three different parties.

| State | Question answered | Where and when | Typical control level |
| --- | --- | --- | --- |
| As Manufactured (AM) | What is this design capable of under controlled conditions? | Test laboratory, before delivery | 0.05 ppm |
| As Installed (AI) | Does it still do that in this room, with this ductwork and this air supply? | Site, after installation, before occupancy | 0.10 ppm |
| As Used (AU) | Does it still do that with the apparatus, the operator and the procedure actually in use? | Site, in service | 0.10 ppm |

3.1 Why AM cannot be replaced by AI

An AM test is run in a test room with controlled cross-drafts, a dedicated exhaust, and no neighbouring fume hoods competing for air. It isolates the hood. That is its value: it is the only test in which a failure is unambiguously attributable to the product.

An AI test includes the room. Supply diffusers, door gaps, adjacent hoods, the position of the make-up air grille and the pressure regime of the laboratory all enter the result. This is why AM results are not transferable to a project and why a manufacturer's type-test certificate cannot substitute for commissioning in a real laboratory.

3.2 Why AU is the only state that reflects risk

AU includes the apparatus. A 20 L reaction flask set on the work surface, a hotplate at the rear of the hood, or a centrifuge tube rack standing 150 mm high in front of the lower baffle slot changes the flow field inside the hood far more than most users expect. The NEBB retrospective attributes 25 % of tracer-gas failures to user practice alone.

AU is therefore the state in which a hood's real protection is decided — and the state most often skipped, because it requires the laboratory to be running the actual process.

3.3 The Deiiang position on test states

Deiiang supplies AM test data for the hood models that have been type-tested, and supports AI commissioning on site through a documented procedure (Section 10). AU testing is a site activity performed with the owner's apparatus; Deiiang's contribution is the control layer that keeps the hood stable while the process changes, described in Section 9.

No manufacturer can certify AU performance, because AU depends on what the owner puts inside the hood. A tender document that requires the supplier to "guarantee AU ≤ 0.10 ppm" is asking for something that cannot be contractually delivered without control of the laboratory's operating practice.

4. Test 1 — Face velocity

Face velocity is the first test in the sequence, and the only one cheap enough to repeat every year.

4.1 Procedure

The sash is set to the specified test opening, conventionally 500 mm (19.7 in) above the work surface for the standard test and additionally at 25 %, 50 % and 100 % of full opening where the specification calls for a velocity profile. The opening is divided into a grid in which no cell exceeds 1 ft² (0.09 m²) and no cell edge exceeds 330 mm (13 in).

A calibrated thermal anemometer is held at the centre of each cell, perpendicular to the opening, and readings are logged for 20 s at 1 s intervals. The 20 s mean is the cell value; the arithmetic mean of the cell values is the hood's face velocity.

Uniformity is reported as the deviation of individual points from the mean. Deiiang reports this as a deviation ratio; the type-tested T3 hood recorded 6.0 %, against a commonly specified ceiling of 20 %.

4.2 What the number means, and what it does not

Face velocity is a screening quantity. It is cheap, fast, repeatable, and it detects the failure modes that dominate field data: clogged filters, slipped fan belts, jammed dampers, leaking ductwork, competing hoods on a shared fan. The Shanghai survey of 27 enterprises found 51.9 % of fans delivering 70 % or less of rated airflow — a fault that face velocity catches immediately.

It does not measure containment. Readings at the geometric centre of a grid cell say nothing about the eddy that forms behind a partially open sash, and nothing about the leakage path around the sash stile. The National Environmental Balancing Bureau's retrospective found that more than 80 % of hoods that failed tracer gas testing had a compliant face velocity. That single figure is the strongest argument in the literature for running both tests.

4.3 Cross-drafts

Before the grid is measured, the standard practice is to measure the cross-draft in front of the hood: probes at approximately 0.5 m (18 in) from the opening, at mid-height, logging for 30 s. SEFA guidance holds cross-drafts below 30 % of the design face velocity. A supply diffuser delivering 2 m/s at 1.5 m from a hood face will destroy containment irrespective of the hood's own quality, and no amount of product engineering compensates for it.

Cross-draft is a room problem, not a hood problem, and it belongs in the AI report as an explicit measured value rather than a note.

4.4 Deiiang velocity data

| Model | Nominal size (W×D×H, mm) | Rated exhaust | Face velocity | Uniformity |
| --- | --- | --- | --- | --- |
| T3 high-performance | 1500 × 850 × 2350 | 1105 m³/h | 0.50 m/s mean | 6.0 % deviation |
| X1 all-steel | 1200 / 1500 / 1800 × 850 × 2350 | 1700 m³/h (1500 mm) | 0.30–0.50 m/s | per project |
| PP hood | 1500 × 850 × 2350 | per configuration | 0.30–0.50 m/s | per project |

The X1 figure is internally consistent and worth showing as a sanity check that specifiers can repeat: 1700 m³/h is 0.472 m³/s, which over the 1500 × 800 mm working opening gives 0.39 m/s — inside the 0.30–0.50 m/s band and below the 0.51 m/s upper bound of the Z9.5 starting range, which leaves headroom for the filter and duct allowance that accumulates over a service life.

A note on the range: 0.30–0.50 m/s is a product range, not a compliance claim.

The applicable acceptance value is whichever the project's specification names — Z9.5's 80–100 fpm starting range equates to 0.41–0.51 m/s, JGJ 91-2019 gives three tiers from 0.35 to 0.75 m/s depending on the toxicity of the material handled, and a low-flow hood designed for 0.30 m/s is legitimately compliant with a specification that permits it and non-compliant with one that does not.

5. Test 2 — Flow visualisation

Before containment is quantified, it is watched: the smoke test is what turns a number into an explanation.

5.1 Procedure

Visualisation is qualitative and it is the test that most often explains a tracer gas failure. Two releases are used. A low-flow release from a TiCl₄ stick or a small smoke source is made just inside the sash plane, around the perimeter of the opening at roughly 150 mm (6 in) inside, to reveal escape at the stiles, the sill and the header. A high-flow release from a smoke generator inside the hood body reveals the internal pattern:

whether the smoke is drawn cleanly into the baffle slots, whether it rolls forward under the sash, and how long the hood takes to clear.

Acceptance is by observation and must be recorded as such: no visible escape at the perimeter, no forward roll-out over the sill, no stagnant region in the working zone, and a clearance time consistent with the design airflow.

5.2 What a good pattern looks like

In a correctly proportioned hood the smoke divides into three streams: the bulk of it leaves through the lower rear slot, a smaller share through the middle, and the lightest fraction, rising buoyantly, through the top slot. That division is not arbitrary. Published design guidance and CFD comparison work converge on an allocation of roughly 50–60 % of exhaust through the lower slot, 20–30 % through the middle and 10–20 % through the top.

The lower slot does the safety-critical work. Dense vapours and spill at the sill are captured there, and it is the slot whose geometry is most often compromised — by a hotplate pushed against the rear baffle, by a waste bottle standing in front of it, or by a baffle that has been adjusted to a horizontal position and never reset.

5.3 Geometry that produces the pattern

The Deiiang hood is built to a three-slot rear baffle with a profiled underside. Four dimensional rules, all traceable to published CFD work, are applied:

| Parameter | Design rule | Source of the rule |
| --- | --- | --- |
| Lower slot allocation | 50–60 % of total exhaust | DAI Scientific CFD comparison, WSU facilities standard |
| Underside baffle height | ≥ 50 mm | Shandong University parametric CFD, 2021 |
| Underside baffle length | ≥ 90 mm | Same study |
| Deflector radius | ≤ 70 mm | Same study — larger radii excite secondary vortices |

The Shandong study is worth understanding rather than merely quoting. It found that when the underside baffle is short and low, a large vortex forms at the centre of the hood and a second vortex forms at the junction between baffle and sidewall. Increasing baffle height past 50 mm and length past 90 mm suppressed both. Adding a deflector eliminated the centre vortex entirely — but only while the deflector radius stayed at or below 70 mm;

beyond that, the deflector's own trailing edge began to generate a vortex that had not existed before. The lesson is that these are thresholds, not monotonic improvements, and that a hood profile which is "more curved" is not automatically better.

Perforation shape matters as well. CFD work validated against experiment at UiTM found circular perforations produced the highest flow uniformity index of the shapes tested, with uniformity improving as perforation count rose. Deiiang's mid-height section uses vertical perforation rather than the single horizontal slot of the 1970s to 1990s legacy pattern.

5.4 Sash and bypass

The sash is 5 mm tempered glass on the X1 and T3 platforms, carried on a cable lift with infinite hold, with PP clamping strip and handle. Tempered glass is specified because its hard surface survives the abrasion of daily wiping and track friction better than any polymer; it fails safe into small blunt fragments, and its transmission stays above 90 % for the service life of the hood.

Two exceptions are worth stating for specifiers. Hydrofluoric acid attacks glass and fogs it permanently, so HF service requires a polycarbonate or PP sash regardless of the abrasion penalty. High-security and high-pressure work calls for laminated glass, whose interlayer keeps fragments bonded and prevents the opening from being breached.

Specifications that follow the 6 mm convention of European practice should state it at order stage; the 5 mm sash on the standard platform satisfies GB 15763.2 and is adequate for normal service, but it is not what a 6 mm clause asks for.

5.5 Smoke test record

The visualisation result should be recorded as a narrative plus a sketch of the pattern at each sash height tested, not as a tick box. A report that says only "smoke test passed" cannot be compared with the next test, and comparison over time is the entire point of routine re-verification.

6. Test 3 — Tracer gas containment

Tracer gas is the only test in the standard that produces a figure comparable between hoods, laboratories and years.

6.1 Apparatus

The tracer is sulphur hexafluoride, released through the ejector defined in the standard's figures. The ejector sits inside the hood, 150 mm (6 in) behind the sash plane, with its calibrated critical orifice passing 4 L/min at 30 psi. A manikin stands in front of the hood with its nose approximately 75 mm (3 in) forward of the sash plane, and the sampling probe is at the breathing zone.

The analyser is an electron-capture or infrared instrument capable of resolving 0.01 ppm, logging at 1 Hz.

SF₆ is used because it is absent from normal laboratory air, roughly five times heavier than air, chemically inert, and therefore unambiguous when it appears at the breathing zone. Those same properties make it an extremely potent greenhouse gas — a global warming potential near 24,000 — which is why some jurisdictions now restrict its use for as-used testing and why alternative tracers are under active discussion. A test programme should note which tracer was used;

the result is not transferable between tracers without correction.

6.2 The three static positions

The manikin is placed at the left, centre and right of the opening in turn, and at each position the concentration is logged for 5 minutes at 1 Hz. The reported values are the 5-minute average, the minimum and the maximum at each position. The hood's containment figure is conventionally the highest of the three position averages; the peak is reported separately and is the value that reveals an intermittent escape event that an average would smooth away.

Three positions is a minimum. A hood that passes left and centre and fails at the right stile has a specific and diagnosable defect — usually sash seal geometry or a sidewall bypass — and a single centre-position test would never find it.

6.3 Sash movement effect

The dynamic test is the one that separates a well-controlled hood from a well-built one. With the manikin at the centre, the sash is cycled through three full opening and closing movements at a controlled speed, while concentration is logged continuously. Two figures are reported: the peak, and the 45 s rolling average (RA-Max).

This test measures the control system as much as the hood. A VAV terminal that responds in 3 s instead of 1 s, an actuator that is under-torqued for the damper, or a controller with rigid PID parameters will produce a spike at the moment the sash moves — a spike that no amount of baffle optimisation can prevent. It is the reason Section 9 treats the control layer as part of the containment system rather than as an accessory.

6.4 Perimeter scan

The optional perimeter scan traverses the sampling probe around the sash perimeter, at the gap between sash and frame, to find localised leakage paths. It is slow and it is the test most likely to find a gasket defect that the three static positions miss.

6.5 Reading the numbers honestly

A reported 0.00 ppm means "below the detection limit of the analyser used", not "zero leakage". Reports should state the analyser's detection limit and the averaging basis. An instrument resolving 0.01 ppm cannot report 0.001 ppm, and a laboratory that compares a 0.00 ppm result from a 0.01 ppm instrument with a 0.003 ppm result from a 0.0001 ppm instrument is comparing instruments, not hoods.

The same caution applies to the difference between an average and a peak. A hood averaging 0.02 ppm with a peak of 0.4 ppm is not obviously safer than one averaging 0.05 ppm with a peak of 0.06 ppm; it depends on whether the peak coincides with a person being present, which it usually does, because the peak is caused by the person moving the sash.

7. Acceptance criteria and the citation trap

The limits are the part of the method most often misquoted, so they are set out here with their sources attached.

7.1 Where the limits come from

| Source | What it supplies | What it does not supply |
| --- | --- | --- |
| ASHRAE 110 | Apparatus, procedure, reporting arithmetic | Any pass/fail value |
| Z9.5-2022 §3.4 | AM 0.05 ppm, AI 0.10 ppm, AU 0.10 ppm | Test procedure detail |
| SEFA 1 | Obligation to publish AM data; cross-draft guidance | Site acceptance |
| Owner specification | The value actually contracted | — |

7.2 The most common errors

Four errors recur in tender documents and acceptance reports, and each one has a consequence.

The first is quoting 0.05 ppm as the site acceptance value. 0.05 ppm is the as-manufactured control level, achieved in a controlled test room. Applying it to a commissioned laboratory rejects hoods that are performing correctly, and it is not what Z9.5 asks for.

The second is treating ASHRAE 110 as a pass/fail standard. It is not one. A report that states "ASHRAE 110: pass" without a control level and a source has not recorded a result.

The third is conflating face velocity with capture velocity. Face velocity is measured across the working opening; capture velocity, under AQ/T 4274-2016 or equivalent, is measured near the source in the workplace. They are different physical quantities with different criteria, and putting both in one compliance table produces a number that means nothing.

The fourth is quoting a single face velocity figure as if it were the design criterion. JGJ 91-2019 sets three tiers by maximum allowable concentration — 0.35 / 0.25, 0.50 / 0.40 and 0.75 / 0.65 m/s for mean and minimum respectively. Judging every hood against a flat 0.5 m/s both over-tests benign work and under-tests highly toxic work.

7.3 A defensible acceptance clause

An acceptance clause that survives audit names five things: the standard and edition; the test state; the control level and the document it comes from; the criterion by which the mean is judged (mean of means, or minimum point); and the instrument's detection limit. Anything less cannot be reproduced by a second party, and a result that cannot be reproduced is not evidence.

8. Deiiang product compliance matrix

This section states what has been measured, on which platform, by whom, and what remains a project-specific obligation.

8.1 What has been tested, and how

| Platform | Construction | Test basis | Key measured results |
| --- | --- | --- | --- |
| T3 high-performance hood | Steel body, corrosion-resistant liner, profiled three-slot baffle | ASHRAE 110-2016, AM, independent test body | 0.50 m/s mean; 6.0 % uniformity; 41 Pa resistance; 0.00 ppm tracer at three positions; RA-Max 0.00 ppm; 1105 m³/h |
| X1 all-steel hood | 1.0 mm galvanised steel, epoxy powder, 5 mm compact liner, 12.7 mm phenolic top | GB/T 24820-2024, full structural and mechanical protocol | Dimensional deviation within ±2 mm; flatness and squareness within tolerance; static load 2000 N; stability under 50 kg impact |
| PP hood | Polypropylene body and liner | GB 24820-2009 | ASS 24 h corrosion grade 7–8; adhesion grade 0; impact 3.92 J undamaged |

The T3 result deserves to be read carefully rather than quoted as a headline. A 41 Pa resistance against a 62.5 Pa ceiling means the hood is not working the fan hard to achieve its containment, which is the difference between a hood that passes on the day of test and one that still passes at year eight when the fan has lost 24 % of its airflow — the retention figure measured for fans at eight years' service.

The two structural standards are not substitutes for ASHRAE 110 and are not presented as such. GB/T 24820-2024 and GB 24820-2009 verify that the furniture is square, strong, stable and correctly finished; they say nothing about airflow. A hood can be structurally perfect and contain nothing.

8.2 Design provisions traced to test items

This is the table a technical evaluator should ask for. Each row states a test requirement, the physical feature that satisfies it, and the consequence of omitting that feature.

| ASHRAE 110 item | Deiiang design provision | Consequence if absent |
| --- | --- | --- |
| Face velocity uniformity | Three-slot baffle with proportional allocation; vertical perforation at mid-height; profiled underside | Local low-velocity zones at the stiles; average looks compliant while a point sits below 60 % of mean |
| Containment at the sill | Lower slot at 2.5 in class width capturing 50–60 % of exhaust; sill airfoil | Dense vapour spill-back onto the operator's forearms |
| Containment under buoyant load | Top slot at 10–20 % allocation | Hot vapours collect under the header and escape on sash movement |
| Vortex suppression | Underside baffle ≥50 mm high, ≥90 mm long; deflector radius ≤70 mm | Centre vortex and baffle-to-sidewall vortex; smoke rolls forward |
| Visual clarity and abrasion life | 5 mm tempered glass sash, cable lift with infinite hold | Scratched or clouded sash; operator raises it higher than needed |
| Dynamic containment | VAV terminal with pressure-independent control, <1 s response, matched actuator torque | Concentration spike at every sash movement |
| Repeatability across service life | Low resistance design (41 Pa measured); accessible baffle; sealed duct joints | Progressive loss of flow; pass at commissioning, fail at year three |
| Perimeter integrity | Continuous PP clamping strip; gasket at sash frame; sealed sidewall penetrations | Localised leak found only by perimeter scan |

8.3 What Deiiang does not claim

Three honest limits belong in any compliance dossier.

AM data exists for the T3 platform; it does not exist for every size and variant, and a project specifying a 1800 mm X1 should ask for the AM data for that width rather than inferring it from the 1500 mm figure. Containment does not scale linearly with width, because the sidewall boundary layer becomes a smaller fraction of the opening as the hood widens.

AI results are site results. Deiiang can supply the procedure, the control layer and the commissioning support, but the number belongs to the laboratory and the test agency.

AU results depend on the apparatus. No supplier can warrant them.

9. The control layer: why a test result is reproducible

A fume hood is a passive device until it is connected to a control system. The containment figures in Section 8 were obtained with the airflow held stable; keeping it stable for the next ten years is the job of the layer described here.

9.1 Terminals

| Device | Function | Relevance to ASHRAE 110 |
| --- | --- | --- |
| Venturi VAV valve | Pressure-independent flow control, ±3 % accuracy, 150–750 Pa independence, <1 s response | Holds face velocity through duct pressure changes from other hoods |
| CAV valve | Fixed flow, pressure-independent, no external power | Stable baseline for constant-volume hoods; no control-induced oscillation |
| Butterfly VAV valve | Fast modulating control with venturi measurement | Cost-effective dynamic response for general chemistry |
| Pitot-tube VAV valve | High accuracy with host communication | Where the specification demands logged airflow |

9.2 Controllers

The FHC10/11/12/100 family drives the terminal from sash position and face velocity. FHC10 integrates controller, sensor input and alarm in one housing with Modbus RTU, an energy-saving mode and emergency exhaust; FHC11 adds an independent communication port and relay outputs for BMS integration; FHC12 accepts third-party actuators, which matters in retrofit; FHC100 accepts displacement or velocity sensors and incorporates the automatic sash motor controller.

The QAVC10/12 and ZPC211 handle room-level flow and pressure: variable or constant volume, room pressure gradient control, noise below 30 dB(A), BACnet and Modbus. They are what keeps a suite of hoods from fighting each other when one sash opens.

9.3 Automatic sash and presence detection

The FSC10 sash controller supports manual, follow and energy-saving modes, with presence detection, foot switch and infrared anti-pinch, and outputs sash position as an analogue signal. Two effects bear on ASHRAE 110 results. The first is direct:

an automatically closing sash holds the working opening at the design height, which is the single most effective containment measure available, because it removes the user behaviour that NEBB attributes to a quarter of all tracer-gas failures. The second is indirect: the analogue position signal lets the VAV terminal anticipate the airflow change instead of reacting to it, which shortens the concentration spike during sash movement.

9.4 Fan and duct

The EC permanent-magnet motor range (0.75–22 kW integrated drive; 11–55 kW direct-drive) holds efficiency in the low-speed range where laboratories actually operate, with 12 protection functions including locked rotor, phase loss and overload. Constant-pressure control through the VFD cabinet keeps duct static pressure flat as hoods open and close, which is the condition under which a VAV terminal's pressure independence is actually tested.

Ductwork is extruded PP from φ110 to φ1000 mm with moulded fittings. Two details affect measured containment more than their cost suggests: branch entries should be radiused rather than square-tapped, and the expansion angle at a transition should be held below 60°, because the Shanghai survey found 46.9 % of exhaust hoods exceeding the 90° limit and paying for it in static pressure and in unstable flow at the hood collar.

10. As-installed commissioning: a twelve-step protocol

This is the sequence Deiiang recommends for AI testing. It is ordered so that each step's precondition is established by the step before it.

- 1. Verify the room before the hood. Measure supply and exhaust airflow, room differential pressure and door undercut. A hood cannot be commissioned in an unbalanced room.

- 2. Confirm the design sash position in writing. Every subsequent measurement is referenced to it, and a later test at a different opening is not comparable.

- 3. Measure cross-draft at 0.5 m in front of the opening, 30 s average. Record it even when it passes.

- 4. Establish airflow at the design value; record fan speed, damper position and duct static pressure.

- 5. Run the face velocity grid at the design sash position, then at 25 %, 50 % and 100 % opening.

- 6. Record the uniformity, not only the mean. Flag any point below the specification's minimum.

- 7. Run low-flow visualisation around the sash perimeter with the laboratory in its normal operating state.

- 8. Run high-flow visualisation inside the hood; record pattern and clearance time.

- 9. Run tracer gas at left, centre and right, 5 minutes each, 1 Hz.

- 10. Run the sash movement test, three cycles, report RA-Max and peak.

- 11. Run the perimeter scan where the specification requires it or where step 7 showed any doubt.

- 12. Post the result. Date, mean face velocity, sash position, test state and any restriction, on the hood itself.

Steps 1 and 2 are the ones most often skipped and most often responsible for a failed commissioning that gets blamed on the product. A hood tested in a room at −40 Pa with a diffuser blowing at its face will fail, and the failure will be recorded against the hood.

11. As-used verification and the annual cycle

Commissioning proves a hood on one day; the annual cycle is what keeps the number true afterwards.

11.1 What routine testing must contain

Z9.5-2022 requires routine performance testing at least annually, and after any significant change to the hood system or the laboratory. The minimum scope is: inspection of hood and room; velocity measurement; cross-draft measurement; VAV response and stability where applicable; verification of presence sensor and controls; verification of monitor and alarm; and airflow visualisation.

Note what is not on the list: the full tracer gas test. Where operation has been demonstrated to be equivalent to the commissioning condition, periodic tracer gas is not mandatory. Where the process or the room has changed, the AI test is repeated instead. That distinction saves laboratories a great deal of money and is routinely ignored in both directions — some sites run tracer gas annually out of habit, others never re-run it after rebuilding the room.

11.2 The 10 % rule

A hood whose average face velocity has fallen more than 10 % below its designated value is to be labelled out of service or restricted use, with corrective action before further work. Two things follow. The designated value must have been recorded at commissioning, which is why step 2 of the protocol exists. And a hood that reads 8 % low is not "fine" — it is a hood with a trend, and the trend is the early warning that the tracer gas test would confirm.

11.3 Signage

Each hood carries the date of its last test, its measured average face velocity, the sash position at which that value applies, and any restriction on work or materials. A hood that has failed carries a restricted-use notice stating the sash position required for safe operation. Signage is not administration; it is the mechanism by which the commissioning number reaches the person standing in front of the hood.

11.4 Degradation to expect

Fan airflow retention falls to roughly 92 % at three years and 76 % at eight years, with corresponding power loss. Filter clogging accounts for 40–50 % of insufficient-airflow faults in cleanroom and laboratory HVAC, belt slack and wear for 20–30 %, and duct leakage for 15–20 %. These are predictable, which means they can be scheduled rather than discovered.

12. Why compliant hoods still leak

The field data is consistent across countries and decades, and it is the best argument for testing containment rather than assuming it.

| Finding | Value | Source |
| --- | --- | --- |
| Hoods with compliant face velocity that exceeded the Z9.5 leakage guideline | 30–50 % | SafeLab, several thousand ASHRAE 110 tests |
| Hoods failing tracer gas that had compliant face velocity | > 80 % | NEBB retrospective |
| Attribution of tracer-gas failure: room conditions / hood design and layout / user practice | 50 / 25 / 25 % | NEBB |
| Finnish survey: compliant face velocity / all criteria met | 27 % / 2.7 % | Pasanen et al., 295 hoods |
| Iranian petrochemical: velocity compliant / tracer compliant, uncorrelated | 9.1 % / 4.5 % | ASHRAE 110 assessment, 22 hoods |
| South African forensic laboratory: outside standard band | 82 % (70 % excessive) | Jacobs, 49 hoods over 11 months |
| Chinese national inspection: overall fume hood non-compliance | 18.7 % | 2025 special inspection bulletin |

The South African figure is the one that most often surprises operators: 70 % of the non-compliant hoods were running too fast, not too slow. Excess face velocity creates turbulence at the sash edge and at the airfoil, and turbulence transports contaminant out of the hood. A hood "improved" by opening its damper to raise the velocity can be made less safe, which is why the acceptance band has an upper bound as well as a lower one.

12.1 Failure modes and the test that catches them

| Symptom | Underlying cause | Test that catches it |
| --- | --- | --- |
| Odour at the operator's face, face velocity normal | Room cross-draft, supply jet at the opening, door gap | Cross-draft measurement, low-flow visualisation |
| Smoke rolls forward over the sill | Lower slot blocked by apparatus, baffle misadjusted | High-flow visualisation |
| Left–right asymmetry | Sash seal, sidewall bypass, uneven baffle | Three-position tracer gas, perimeter scan |
| Spike whenever the sash moves | VAV response lag, actuator torque, PID tuning | Sash movement test, RA-Max |
| Gradual loss of containment over years | Filter loading, fan belt wear, duct leakage | Trend of face velocity against designated value |
| Intermittent alarm, normal readings otherwise | Sensor drift, probe in an eddy | Sensor calibration, cross-check against reference |

12.2 Detection lag

Faults in laboratory ventilation are typically detected an average of 72 hours after they begin, because most systems have no online fan-status monitoring and no trend logging. A hood with a face velocity sensor and a BMS point turns a three-day exposure into a same-day alarm. This is the strongest practical argument for the control layer in Section 9 that has nothing to do with energy.

13. What a defensible test report contains

A report that will survive a review by a second test body contains the following. Anything missing is a gap a reviewer will find.

| Element | Why it matters |
| --- | --- |
| Standard and edition | 1995 and 2016 results are not interchangeable |
| Test state (AM / AI / AU) | Determines which control level applies |
| Control level and its source | Otherwise the result has no pass/fail meaning |
| Room conditions at test: temperature, pressure differential, supply and exhaust airflow | The 50 % room attribution is invisible without this |
| Cross-draft values and measurement positions | Same |
| Sash position for every velocity measurement | A velocity without a position is not reproducible |
| Complete velocity grid, not only the mean | Hides the low-velocity point that matters |
| Visualisation narrative and pattern sketch | Comparable over time; a tick box is not |
| Tracer gas: tracer identity, release rate, ejector position, manikin position, analyser model and detection limit | Determines whether 0.00 means anything |
| Per-position average, minimum and maximum | The peak reveals intermittent escape |
| Sash movement: cycle count, speed, RA-Max, peak | Tests the control system, not the box |
| Instrument calibration certificates and dates | Without them the numbers are opinions |
| Test body accreditation | Determines evidential weight |

14. Limitations and honest notes

Three limitations belong in the document rather than in a footnote.

ASHRAE 110 does not set a safe exposure level. A hood achieving 0.05 ppm AM is a hood that contains well under controlled conditions. Whether 0.05 ppm of a specific substance for a specific duration is acceptable is an industrial hygiene question that the test cannot answer, and SEFA states this explicitly in its AM clause.

A test is a snapshot of one day. Temperature, room pressure, duct condition, neighbouring hood operation and the presence of apparatus all change the result. The value of a test programme lies in repetition against a recorded baseline, not in a single certificate.

Tracer choice is changing. SF₆'s greenhouse warming potential is driving restriction in some jurisdictions, particularly for as-used tests. Projects in those markets should state the permitted tracer in the specification, because a result obtained with SF₆ is not directly comparable with one obtained with a substitute without a correction that the standard does not currently provide.

15. References

- 1. ANSI/ASHRAE 110-2016, Method of Testing Performance of Laboratory Fume Hoods. ASHRAE, Atlanta.

- 2. ANSI/ASHRAE/ASSP Z9.5-2022, Laboratory Ventilation. Control levels §3.4; design face velocity §4.3.1; routine testing §7.3.

- 3. SEFA 1, Laboratory Fume Hoods — Recommended Practices. Scientific Equipment and Furniture Association.

- 4. EN 14175-2:2003 and EN 14175-3:2006, Fume Cupboards. CEN.

- 5. JGJ 91-2019, Code for Design of Scientific Research Buildings, Table 8.3.2. China.

- 6. GB/T 24820-2024, General Technical Requirements for Laboratory Furniture. China.

- 7. Yan R., Dong Z., Lu S., Lai Y., Lü M. Analysis of the influence of the underside baffle and deflector of the fume hood on the flow field. Journal of Shandong University (Engineering Science), 2021, 51(5): 122–130.

- 8. Omar M.A.A. Effect of Baffle Design Configuration to the Flow Performance of a Servco Fume Cupboard. MSc thesis, UiTM, 2018.

- 9. DAI Scientific / Flow Sciences. Bypass Fume Hoods Compared Using CFD — Legacy vs STF Design, 2023.

- 10. Pasanen T., Kalliokoski P., Kokotti H. Performance of Laboratory Fume Hoods. University of Kuopio; AIVC.

- 11. Evaluation of laboratory fume-hood performance in a petrochemical industry based on ASHRAE 110. Iran.

- 12. Jacobs P.J. Laboratory Fume Hood Performance. MSc thesis, University of the Witwatersrand, 2009.

- 13. NEBB. The Myth of Face Velocity Reloaded. National Environmental Balancing Bureau.

- 14. SafeLab. Using the ASHRAE 110 Test as a TQM Tool to Improve Laboratory Fume Hood Performance.

- 15. 2025 National Laboratory Safety Hazard Special Inspection Bulletin. China.

- 16. Survey of Local Exhaust Ventilation Facilities in 27 Manufacturing Enterprises, Shanghai. Journal of Occupational Health and Emergency Rescue, 2024(5).

- 17. Crowe E. et al. How Faulty are Commercial Buildings? LBNL / PNNL, 2023.

- 18. Deiiang type-test data sheet, T3 fume hood, ASHRAE 110-2016 (AM). Independent test body, 2025.

- 19. Deiiang test report BT2504G1590554GR, GB/T 24820-2024, X1 all-steel fume hood. 2025.

Appendix A — As-installed commissioning data sheet

| Field | Entry |
| --- | --- |
| Project / laboratory |  |
| Hood model and serial |  |
| Nominal size (W×D×H) |  |
| Test standard and edition |  |
| Test state | AI |
| Control level applied and source | 0.10 ppm, Z9.5-2022 §3.4 |
| Room temperature / pressure differential |  |
| Supply airflow / exhaust airflow |  |
| Cross-draft, mean of 30 s |  |
| Design sash position |  |
| Face velocity, mean |  |
| Face velocity, minimum point |  |
| Uniformity deviation |  |
| Visualisation, low flow |  |
| Visualisation, high flow and clearance time |  |
| Tracer gas, left / centre / right average |  |
| Tracer gas, maximum at any position |  |
| Sash movement RA-Max / peak |  |
| Perimeter scan result |  |
| Analyser model and detection limit |  |
| Instrument calibration dates |  |
| Test body and accreditation |  |
| Result and any restriction |  |
| Tester signature and date |  |

Appendix B — Abbreviations

| Term | Meaning |
| --- | --- |
| AM / AI / AU | As manufactured / as installed / as used |
| CAV / VAV | Constant air volume / variable air volume |
| LEV | Local exhaust ventilation |
| RA-Max | Maximum 45 s rolling average, sash movement test |
| SF₆ | Sulphur hexafluoride, tracer gas |
| TiCl₄ | Titanium tetrachloride, visualisation smoke source |
| EC motor | Electronically commutated permanent-magnet motor |
| fpm | Feet per minute; 100 fpm = 0.51 m/s |